Super Pulsed Regimes of Diode and Fiber Lasers for Surgical and Therapeutic Medical Applications

Ilya Yaroslavsky1, Victoria Andreeva2, Alexander Vybornov1, Igor Perchuk1, Gregory Altshuler1

1IPG Medical, Marlborough, MA; 2 IRE POLUS, Fryazino, RU

Recent advances in laser technology expended applications of lasers in many medical fields. Ability to deliver sufficiently high energy in a short pulse while maintaining required average power (“super pulse”) is advantageous for a number of surgical and therapeutic procedures. Super Pulse (SP) regimes in Laser Diodes (LD) and Fiber Lasers (FL) offer numerous benefits over existing techniques. In this work, we have conducted thorough evaluation of two super-pulse LD and FL systems, respectively designed for treatment of soft tissues (in surgical and therapeutic modes) and stones of urinary tract (lithotripsy).

# Systems evaliated:

- Super pulse LD system providing up to 30 W average and up to 150 W peak power at 980 nm wavelength. The laser was operated in various modes of control including those with thermal feedback. The following characteristics of the system performance were evaluated: 1) Speed and depth of cutting; 2) Degree of charring and dimensions of coagulative margin; 3) Effective treatment depth in therapeutic mode;

- Super pulse Thulium (Tm) FL system for lithotripsy capable of operation with up to 50 W average and 500 W peak power at 1940 nm wavelength. The in vitro experimental setup included a specially designed cuvette allowing quantitative assessment of size distribution of stone fragments and evaluation of the ablation rate. Magnitude of the retropulsion effect was measured as well.

# Results:

- The SP laser diode system was compared with industry-leading conventional diode and CO2 devices. The results indicated that the super-pulse diode laser system yielded increase in speed of controlled cutting by a factor of 2-3 in comparison with the conventional diode laser and approaching that of CO2 device;

- In therapeutic mode, the SP regime of the laser diode system provided substantial increase in the effective treatment depth compared with CW regime at equal average power;

- The SP Tm fiber laser system was compared with leading Ho:YAG system on the market. The two systems were matched in terms of average power and/or pulse energy. Ablation rate measurements revealed that, depending on stone composition and laser settings, the Tm laser provided between 1.2 and 4.3 times faster ablation than the Ho system. The average retropulsion distance after a single pulse was substantially shorter with the Tm laser.

# Conclusion:

Super-pulse technology implemented in laser diodes and fiber lasers may have significant potential for creating a new standard of care in surgical and therapeutic medical applications.